## Slide 1: Pathlength-selective, interferometric Speckle Contrast Optical Spectroscopy
Non-Invasive Methods for Brain Research Pathlength-Selective Interferometric Speckle Contrast Optical Spectroscopy
Shira Tziony
Supervisors: Dr. Nisan Ozana and Ms. Marina Vasileva
July 28, 2026

## Slide 2: Table of content
Brain Imaging: Motivation and Limitations
Uses for Brain Imaging
Advantages and Limitations
Chosen method- Speckle Contrast Optical Spectroscopy
Rational
Main principles
Introducing Time and Coherence Gating
Temporal Point Spread Function and Instrument Response Function
Experiments
Setup
Noises and Limitations
Metrics
Pipeline and GUI
Results
TPSF and IRF measurements
Signal-to-Noise Ratio
Beats Per Minute
Sensitivity
Discussions and Conclusions
TPSF and IRF Performance
SNR and BPM Performance
Sensitivity Metric Limitations
Conclusions and Future Work
Signal Quality

## Slide 3
Brain Imaging –Uses in the Medical Fiels
Stroke diagnosis and cerebral blood flow assessment
Brain tumor detection and surgical planning
Traumatic brain injury (TBI) monitoring
Epilepsy diagnosis and seizure localization
Neurodegenerative disease assessment (e.g., Alzheimer's, Parkinson’s)
Neurocritical care monitoring (ICP, oxygenation, brain function)
Functional brain mapping (motor, language, cognition)
Psychiatric and cognitive disorder research Rehabilitation
Brain-computer interfaces (BCIs)

## Slide 4
Brain Imaging – Non-Invasive Methods Available
For Speckle Contrast Optical Spectroscopy (SCOS):
  Special resolution of a few mm to a few cm
  High temporal resolution, but 50ms is sufficient (20 FPS)
Source: Thompson Rivers University (n.d.)
Source: Zhu, He & Chen (2018)

## Slide 5
The Challenges of traditional SCOS
Sensetivity Limitation: dominated by superficial layers (scalp and skull)
Signal limitation: Increasing source detector separation (SDS) reduces signal quality
The Goal: Enhance depth sensitivity and signal quality by using time and coherence gating
Source: Stanford CIBSR (n.d.)

## Slide 6: Relation Between Contrast to Movement -Rational
Speckle: scattered light randomly self interferes
Disturbance: changes the speckle formation
Integration: speckle movement causes blur
Low contrast -> High movement!
Contrast
Camera Exposure
Signal
Excitation Trigger

## Slide 7: SCOS– Main Principles
Siegert Relations
Contrast squared
Std of an intensity window
Coherence Factor
Exposure Time Limitations
Intensity autocorrelation function
Field autocorrelation function
Average of an intensity window

## Slide 8: SCOS– Main Principles
For our medium:
We can put into the last equation and get:
Speckle decorrelation time
Blood Flow index:

## Slide 9: Pathlength Selective Interferometric SCOS (PaLS-iSCOS)– Introducing Time and Coherence Gating
A coherent Gaussian pulse is split
Sample widens the photon arrival time
Delayed reference amplifies late-arriving photons
A gated sensor can help filter early arriving photons
Reference
Sample
Instrument Response Function (IRF)
Temporal Point Spread Function (TPSF)
Changing distance

## Slide 10: TPSF vs Reference IRF
IRF
Defined by the laser pulse
Defines the coherence gate length:
TPSF
Defines the arrival of photons from the tissue
Later arriving photons will be at the tail
FWHM of laser pulse
Velocity in tissue
Timing of the reference is crucial!
Source: Robinson et al. (2025)

## Slide 11
Revised theory for Time and Coherence Gating
11
Pathlength difference
Coherence length
Effective path length distribution
Time gate
Total Intensity
Gate Profile
SCOS term
Interferometric term

## Slide 12: Experiment Setup and Design
Setup
Noises and Limitations
Metrics
Pipeline and GUI

## Slide 13: PaLS-iSCOS vs Gated SCOS Experiment Setup
Camera: SPAD512 from PI-Imaging
SPAD = single-photon avalanche diode
Laser Source: 765nm laser pulse (FWHM ~350ps, 20MHz)
Probes: 1cm, 1.5cm,2.5cm source-detector separation (SDS)
Laser Power | 
% | mW
30 | 3
35 | 6
40 | 12
45 | 15
55 | 27

## Slide 14: TPSF Measurement With Long Gate
Challenge: camera gate (6ns) is wider than IRF/TPSF FWHM (350ps to 700ps)
Solution: Richardson-Lucy Algorithm, an iterative solution to compute the TPSF
Stopping point: minimum MSE
True TPSF/IRF
SPAD gate (smearing effect)
Image captured by SPAD

## Slide 15: SCOS Noise Model
Doesn’t exist in SPAD
- The measured contrast
- Shot noise
- Read-out noise
- Dark current noise
- Quantization noise

## Slide 16: Added SPAD Noise
-Intensity at one pixel (Dark)
-Standard deviation in pixel intensities in the image Region of Interest (ROI)
Hot pixels threshold

## Slide 17: Metrics
Signal to Noise Ratio (SNR)
Measures the quality of the signal
Sound – amplitude of peak frequency in 0.5Hz to 2.5 Hz
Noise – median amplitude of frequencies in 0.5Hz to 2.5 Hz
BPM Accuracy
Measures the accuracy of the BPM detection with no reference
Additional measure of stability using Coefficient of Variation
Mean and variance are of the detected BPM
Sensitivity
Measures the sensitivity for the superficial layer (scalp)
Baseline- control rBFI measurements
Occlusion- tourniquet of the scalp

## Slide 18: PaLS-iSCOS vs Gated SCOS SNR and BPM Experiments
Coherence gates limited by maximum reference arm length
Source Detector Separation (SDS) | Gated SCOS time gates | PaLS-iSCOS |  | Laser power Level | 
 |  | Time gate | Coherence gate |  | 
 |  |  |  | Gated SCOS | PaLS-iSCOS
1cm | -300ps to 1100ps | -300ps | -50ps to 900ps | 30%, 40% | 30%
1.5cm |  |  | -200ps to 750ps | 45%, 55% | 45%
2.5cm |  |  | -500ps to 400ps | 55% | 55%

## Slide 19: PaLS-iSCOS vs Gated SCOS Sensitivity Experiments
Fewer coherence gates due to time issues
3 stages (30s each):
Baseline
Tourniquet
Rest (only for 1.5cm)
Source Detector Separation (SDS) | Gated SCOS time gates | iSCOS |  | PaLS-iSCOS |  | Laser power Level Gated SCOS PaLS-iSCOS
 |  | Time gate | Coherence gate | Time gate | Coherence gate | 
1cm | 0ps, 250ps, 500ps, 750ps | -300ps | 0ps, 250ps, 500ps | -200ps of the coherence gate | 0ps, 250ps, 500ps | 35%
1.5cm | 0ps, 250ps, 500ps |  |  |  |  | 45%

## Slide 20: PaLS-iSCOS vs Gated SCOS Experiment Pipeline
Measurements and calculations were handled through a custom camera GUI
TPSF measurement
Gated SCOS measurements
IRF measurement and adjustment
PaLS-iSCOS measurement
Measure TPSF and extract the TPSF peak
Measure SCOS and subtract noise
Calculate time difference between peaks
Measure SCOS and subtract noise
Sweep around the TPSF peak for different gates
Measure IRF and extract the IRF peak
Repeat for different pathlengths

## Slide 21
Automatic communication with camera and image taking
Automatic noise and deconvolution calculation
TPSF\IRF, SCOS, and Live intensity windows
GUI Build
TPSF view
IRF view
Live view
SCOS view

## Slide 22: Results
TPSF and IRF Measurements
Gated-SCOS Measurements
PaLS-iSCOS Measurements
SNR
BPM
Sensitivity

## Slide 23: Results –TPSF and IRF
Both TPSF and IRF were measured successfully
FWHM of IRF matches the theoretical one (353.4ps vs ~350ps)
FWHMs make sense

## Slide 24: Results –Gated SCOS (1cm SDS)
Laser power:40%
Early gate has a good signal shape
Late gates are much noisier
SNR matches accordingly
Gate offset -100ps, BPM 73, SNR 18.59dB
Gate offset 200ps, BPM 73, SNR 20.83dB
Gate offset 700ps, BPM 92, SNR 7.92dB

## Slide 25: Results –PaLS-iSCOS (1cm SDS)
Laser power:30%
All with time gates of -300ps, different coherence gates
Early gates show the best SNR, as expected
Late gates show potential as well
Gate offset -74ps, BPM 78, SNR 25.75dB
Gate offset 242ps, BPM 73, SNR 20.90dB
Gate offset 763ps, BPM 82, SNR 9.35dB

## Slide 26: Early Results –SNR (1cm SDS)
Laser Power Level: 30%
Subtracting Noise shows limited improvement
PaLS-iSCOS has improved SNR
Raw Contrast
Corrected Contrast

## Slide 27: Early Results –SNR (1.5cm SDS)
Laser Power Level: 45%
Noise subtraction does not improve signal
PaLS-iSCOS has improved SNR, but for less gates
Raw Contrast
Corrected Contrast

## Slide 28: Early Results –SNR (2.5cm SDS)
Laser Power Level: 55%
Noise subtraction degrades signal
PaLS-iSCOS has lower SNR
Raw Contrast
Corrected Contrast

## Slide 29: Results –BPM Accuracy
Accuracy- between 60 to 100 BPM
1cm SDS
1.5cm SDS
2.5cm SDS

## Slide 30: Results –BPM CV
CV- stability
1cm SDS
1.5cm SDS
2.5cm SDS

## Slide 31: Selected SDS for Sensitivity
1cm SDS:
  Better or similar SNR
  Better or similar signal quality
  Improved to similar accuracy
  Similar stability
1.5cm SDS was measured as well
Gate offset 242ps, BPM 73, SNR 20.90dB
SNR
Stability
Accuracy

## Slide 32: Sensitivity Test Example(1cm)
Laser power:35%
PaLS-iSCOS examples from early and late gate
rBFI normalized based on the base window
The rBFI drop is lower for the late gate
The signal quality is better in occlusion
Normalized rBFI vs Time (-25ps Coherence Gate)
Normalized rBFI vs Time (500ps Coherence Gate)

## Slide 33: Sensitivity Test(1cm)
Laser power:35%
PaLS-iSCOS shows the best sensitivity
Gated SCOS results after 300ps are unreliable

## Slide 34: Sensitivity Test(1.5cm)
Laser power:45%
Noisy results all around

## Slide 35: Discussions
TPSF and IRF Performance
SNR and BPM Performance
Sensitivity Metric Limitations
Signal Quality

## Slide 36: Discussion –TPSF and IRF
While:
  FWHM of IRF matches the theoretical one
  FWHMs make sense
More validations are needed

## Slide 37: Discussion –SNR and BPM Performance
SNR predicts BPM accuracy up to a noise floor (~10dB)
Measurements with ground truth BPM could improve accuracy validations
1cm SDS

## Slide 38: Discussion – Signal Quality
For late gate, poor signal quality
Achieving >10dB late gate could be done by:
  Increasing photon count
  Testing different sample-to-reference ratios
  Using a better camera sensor
Overall, the signal does show the physiological process of bloodflow
Time gate offset 700ps, BPM 92, SNR 7.92dB
Coherence gate offset 763ps, BPM 82, SNR 9.35dB

## Slide 39: Discussion –Sensitivity Metric Limitations
Sensitivity metric is very unreliable
Possible cause:
  Tourniquet fastens/changes probe position
  Intensity spikes
  At low photon count, it affects contrast
Rethinking the sensitivity test or metric is needed
Low Photon Count
High Photon Count

## Slide 40: Conclusions
SCOS relies on speckle contrast to extract rBFi
The method can have good resolution, but lacks sensitivity and accuracy
Adding coherence and time gating could improve both sensitivity and SNR
Results show both improvement of sensitivity and SNR when selecting the right configuration
What is next
Validate TPSF and IRF profiles
Improve signal quality
Improve sensitivity test
Experiment with different sensors and setups

## Slide 41: Thank you for listening! Any Questions?

## Slide 42: References
Zhu, J., He, X., & Chen, Z. (2018). Perspective: Current challenges and solutions of Doppler optical coherence tomography and angiography for neuroimaging. APL Photonics, 3(12), 120902. https://doi.org/10.1063/1.5040590
Thompson Rivers University. (n.d.). Research methods in psycholinguistics. Psychology of Language Pressbooks. https://psychologyoflanguage.pressbooks.tru.ca/chapter/research-methods-in-psycholinguistics/
Stanford Medicine. (n.d.). About NIRS. Center for Interdisciplinary Brain Sciences Research. https://med.stanford.edu/cibsr/GettingReady/about-nirs.html
Robinson, M. B., et al. (2025). Pathlength-selective, interferometric diffuse correlation spectroscopy. IEEE Journal of Selected Topics in Quantum Electronics, 31(4), 1–14. https://doi.org/10.1109/JSTQE.2025.3575719

## Slide 43: Supplementary Slides

## Slide 44: TPSF Measurement With Long Gate
Challenge: camera gate (6ns) is wider than IRF/TPSF FWHM (350ps to 700ps)
Solution: Richardson-Lucy Algorithm, an iterative solution to compute the TPSF
Stopping point: minimum MSE
True TPSF/IRF
SPAD gate (smearing effect)
Image captured by SPAD

## Slide 45: TPSF Measurement With Long Gate
Challenge: camera gate (6ns) is wider than IRF/TPSF FWHM (350ps to 700ps)
Solution: Richardson-Lucy Algorithm, an iterative solution to compute the TPSF
Stopping point: minimum MSE

## Slide 46
Theory behind PaLS-iDCS
As for the (final) autocorrelation function, we get:
where M is the number of detected spatial modes, is the average total intensity that falls within the selected time gate centered at , are the average sample and reference arm intensity respectively, is the length of the reference arm pathlength, and is the polarization-dependent coherence factor:

## Slide 47
Theory behind PaLS-iDCS
Let’s divide the function to two parts (excluding the 1 and ):
1 -
2 -
The first term can be identified as a traditional TD-DCS, While the second term is unique for PaLS-iDCS.
Now, let’s assume a high reference intensity is selected (

## Slide 48
Theory behind PaLS-iDCS
We will drop the first term, giving us:
We still have dependency, but we can subvert that by a few principles that move the dependency to .

## Slide 49
Theory behind PaLS-iDCS
First, if we look at the coherence gate function, we can see that we achieve a peek at :
This function defines the interference window. interference is only significant when the optical pathlength s of the scattered sample photon matches the reference arm pathlength ​ within the coherence length ​ of the light source.
i.e., defines the coherence gate.

## Slide 50
Theory behind PaLS-iDCS
Now, we must take into account and .
While the equation includes dependencies on both the effective pathlength distribution and the instrument response function , in PaLS-iDCS these are not independently varied or scanned.
Instead, due to the interference mechanism, the reference arm length ​ effectively sets the gate center ​, because photons with pathlength ​ will arrive at:

## Slide 51
Theory behind PaLS-iDCS
This means that ​ is chosen based on ​to select photons of specific pathlengths.
This is the connection between time and coherence gate!
Now, let’s discuss how to extract the optical properties of the system.

## Slide 52
Theory behind PaLS-iDCS
Since we are still using the Siegert relation, the general form is still:
So, to infer , we can use and multiple to get at different time points:
From here, we can fit and estimate the optical properties.

## Slide 53
Supp References
Wang, Q.; Bi, R.; Zheng, S.; Erdogan, A.T.; Qi, Y.; Li, C.; Hua, Y.; Pan, M.; Wang, Y.; Finlayson, N.; et al. Fiber-Based Ultra-High-Speed Diffuse Speckle Contrast Analysis System for Deep Blood Flow Sensing Using a Large SPAD Camera. Biosensors 2025, 15, 514. https://doi.org/10.3390/bios15080514